## Scanning
This is the setup for the 78um 30 Direction DTI scan:
1. For the DTI scanning we do not use the corner coils as the SNR is not high enough in those coils.
2. Make sure the ethernet cord is plugged in for channels 10-16.
3. Log into Jacob's account.
  1. Username - jacob
  2. Password - jacob
4. Run spulsmice as you would normally, set up for 16 channel mode.
5. Find the sequence titled "fse3dmice_lsn_diff_30Dir_78um" and double click on it.
6. Go to the scan tab.
7. Currently in my sequence it will load up with an "ERROR See Text Output for More Detail" - For some reason it is appending an extra 16 to the variable RFCOIL and rfcoil. They are incorrectly listed as solenoid_1616 and need to be changed
  1. rfcoil="solenoid_16"
  2. RFCOIL="solenoid_16"
8. Go to grof in the dropdown menu and proceed like normal.
9. Go to phase in the dropdown menu and proceed like normal.
10. Go to gain in the dropdown menu and proceed like normal. Gain for the diffusion sequences needs to be set to ~20 as the initial TE is longer than the 40um sequence.
11. Go to scan.
  1. Change nfmod=6
  2. Run the macro "diff_jacon35inv_new" - This populates the array variables for the diffusion directions.
  3. check dro? Should be an array of 35 variables
  4. Run "mmgo('check')"
  5. Check nf? Should be 60
  6. Check ni? Should be 17
  7. Check nfid? should be 25
  8. Run the macro "time" Should be 57 sec.
  9. Run mmgo - there might be a slight delay prior to it starting while it is populating the arrays
  10. Data will be saved to /hpf/largeprojects/MICe/jacob/fid/ - please copy it over to your own directory and let me know so I can delete it after.